NanoPro 1000: User Guidelines and General Protocol

Sample Preparation

- While making your samples, please ensure all the PBS/media is aspirated out before you add lysis buffer.

- Use a final concentration of 0.1mg/ml – Lysis Buffer Bicine/CHAPS (Protein Signaling Inc.) or MPER (Thermo Scientific Inc.)

- Premix in NanoPro compares to a gel in Western Blot (WB) – classified by ampholyte gradients: (4-9) – broader pI gradient: use this as a starting point unless you know more about the peaks you can expect. / (5-8) – shorter range; also you can use a combination of the two.

- Standards are fluorescently labeled standard pIs, which compare to the standard protein ladder on the WB

- Sample dilutant compares to lysis buffer for WBs

- For a 10ul/well run, make 40uls of solutions for triplicates and 30uls for duplicates – to cater for pipeting losses

Setting up the assay plate

- Make stock cocktails

  - HNG dilutant (20ul/tube)+ Inhibitor cocktail

  - Premix + Standards (3% of total volume) --- 100ul + 3ul --- VORTEX v.well

Examples:     Triplicate (40ul)    Duplicate (30ul)

Premix 		20		15

HNG Dilutant	16.3		11.7 ---- concentration depends on concentration of available sample

Sample 		3.7		8.3 ---- concentration depends on concentration of available sample

- Pipet the samples first into individual tubes

- Add the mix of dilutant and inhibitor in the required concentrations based on the number of samples in the run

- Add the Premix and Standards to add up to 40 or 30 uls (depending on duplicates/triplicates)

- Mix each sample tube with a larger pipet tip to ensure complete mixing – Be very thorough with this step.

- Load the diluted (1:200, or as reqd.) antibody to each of the wells, as below

- Load the sample into the first row

- Sample Plate:  e.g. 4 samples + 1 positive control; 3 antibodies (AB)

| A | 1 Sample 1 | 2 Sample 2 | 3 Sample 3 | 4 Sample 4 | 5 +ve cntrl | Typical examples of antibodies |
| --- | --- | --- | --- | --- | --- | --- |
| B | Housekeeping | Housekeeping | Housekeeping | Housekeeping | Housekeeping | Hsp70 |
| C | AB1 | AB1 | AB1 | AB1 | AB1 | ERK |
| D | AB2 | AB2 | AB2 | AB2 | AB2 | pMEK |
| E | AB3 | AB3 | AB3 | AB3 | AB3 | MEK |
| F | Sec AB | Sec AB | Sec AB | Sec AB | Sec AB | Anti- Rab |
| G | Sec AB | Sec AB | Sec AB | Sec AB | Sec AB | Anti- Mou |
| H |  |  |  |  |  | Luminal Peroxide |

- Spin plate for 5min at 2500 rpm

Setting up machine for the run

- Compass software ( Instrument ( Autoclean (takes 10min)

- Open trays – click Resources

  - Acid – red; Base – Blue; TBS-t – green

  - Refill the containers of water and waste

- Load a new capillary box

- Click Sample on the Open Tray to load your new plate (leave the lid on)

Saving assay template

- File ( New

- Enter the details of the Pr. and Sec. Antibodies and Samples

- Save Assay in desired folder

- Advanced Protocol ( maintain defaults unless otherwise recommended

  - Power – compensates w.r.t. the protein size ---SELECT as reqd.

  - Voltage – constant voltage; use if amounts of protein in each capillary vary

- Immobilization UV exposure = 80s

- Pr. Ab = 120min, Sec. Ab = 20mins; Multiple Exposure time = 30/60/120/700/1000 secs

- Cycle Runs:  A              –             B              –           G                     –             H

Sample      ---           Hsp70       ---        Anti-Mouse       ---           LP

A             –             C              –           F                     –             H

Sample      ---           ERK       ---        Anti-Rabbit          ---           LP

- (Note machine labels the row A1, but the machine picks up all samples in row A.)

- Edit ( Schedule ( Overlap with hold (most effective)

- While the machine is running, click “protocol” to change the protocol.  After making changes, save the protocol.

Contact for further information:

Liwen Xu

CCSR Rm 0128

Phone: 3-5050

Email: xuliwen4@stanford.edu